A LOW COST, HIGH SPEED, 16 CHANNEL LOGIC ANALYZER

This project is a high speed logic analyzer you can build for about two

hundred and fifty dollars.  Here are the key features:

- PC hosted - your current PC provides the user interface.

- 20 or 25 Mhz maximum acquisition rate.

- 6 clock selections (max speed, 1/2, 1/4, 1/8, 1/16, external).

- 16 channels.

- 8 bit trigger word with don't care.

- 8K word buffer - 4K pretrigger and 4K posttrigger.

- easy to learn software and diagnostics.

DESIGN AND THEORY OF OPERATION

The analyzer is made up of two separate modules, an interface board that

resides in a slot in your PC and a main analyzer module external to the

PC.  They are connected by a 50 pin ribbon cable.  Although everything

could have been put on a single board, the modular approach was used for

two reasons.  The first was to keep the data probe leads short as

possible but allow room to work (I felt that a tangle of 18 inches probe

leads coming from the back of a PC would be a bit awkward to work

with.).  The second reason was the save PC slots, I have other projects

in the works that could use the same PC interface card.  By keeping

things modular I only need to use one PC slot and can connect other

units to the interface card as needed.

The interface board is a very standard two chip parallel interface.  The

8255 has three 8 bit parallel ports that can be configured by software.

This provides for a flexible interface.  In normal operation of the

analyzer port A and port C will be configure to be output ports and port

B is configure as an input port.  The second chip is a 74LS682 8 bit

comparator.  It is used as an address decoder for the 8255.  The 74LS682

asserts pin 19 low when it's P inputs are the same as it's Q inputs.

The eight Q inputs have an internal 10 K pull up resistor so they will

stay high unless forced low.  The 8255 address is selected with the DIP

switch, each switch that is turned on will pull the corresponding Q

input low.  When the address on the PC bus (the P inputs) matches the

address selected by the DIP switch (Q inputs) a chip select is generated

for the 8255.  If it is an I/O cycle then either IOR* or IOW* will also

be asserted and data will be read from or written to the 8255 register

selected by A0 and A1.  The 8255 base address is specified to the

software in the configuration file.

50 pin connector P1 is the bridge to the main analyzer board.  Each of

the three 8 bit ports from the 8255 are connected to P1.  Of the

remaining 26 lines, 12 are connected to PC ground, 12 to PC +5 volts

supply, 1 to the PC +12 volts and one to -12 volts.  The +/- 12 volt

supplies are not used by the analyzer but are available for other

projects.

While the interface board design was simple and straightforward, the

design of the logic analyzer board was complicated by the high (20-30

Mhz) clock rates on the board.  Covering the problems of high speed

logic system design is beyond the scope of this article, however a brief

look at two areas of high speed design and how they affected the final

design is warranted.

A key area in high speed logic system design is power distribution.  The

signal transitions in a high speed system occur very quickly (1-2 ns),

for each transition the output drivers of a device require a surge of

power.  These power surges require a low impedance path for power and

ground to each device and to the bypass capacitors.  The typical

solution to this requirement is the use of a multi-layer PC board with

dedicated power and ground planes and carefully placed bypass

capacitors.  The best placement of the bypass capacitors is right at the

device pins such as the bypass capacitors built into the IC socket or

the flat capacitors that are installed under the IC in the same mounting

holes.  However a four layer PC board costs about twice as much as a

double sided board and the initial tooling charges can be about four

times that of a double sided board (the up front costs are very

significant if quantities are small or unknown).  To keep the analyzer

cost low I needed to use a double sided PC board but get the power

distribution quality of a four layer board.  The solution to this

problem turned out to be the use of three QPACs on a double sided board.

QPACs are long capacitors with multiple leads attached to each plate of

the capacitor.  The QPACs used in the logic analyzer are 7.35 inches

long with the leads spaced for six 20 pin DIP ICs.  The QPAC lies flat

on the PC board underneath each IC, this give the effect of having

dedicated power and ground planes and perfectly placed bypass capacitors

for each IC in the row.

Another area of concern for high speed logic system design is signal

routing.  Avoiding sharp bends and board vias in high speed lines

reduces reflections making for cleaner signal transitions.  Seven

GAL16V8 programmable logic devices are used in the analyzer, each of

these replace two or three standard ICs, thus reducing the number of

interconnecting signals, making routing easier.  In addition the use of

programmable logic allowed the rearrangement of device pinouts to make

routing easier.  Because of this there are no vias on any of the high

speed signal lines.

The main analyzer board has four main functional blocks, the PC

interface, clock generation, data storage and trigger control.

The PC interface consists of P1, U5, U6, U7, U8, U9 and U10.  The three

eight bit ports from the 8255 on the interface board are bought onto the

analyzer board through P1.  The three ports are expanded into two output

ports (U9 and U10) and four input ports (U5, U6, U7 and U8).  The two

output ports are driven by 8255 port A and controlled by port C.  Data

is written to port A and then strobed into either U9 or U10 as selected

by port C.  The four input ports U5, U6, U7 and U8 work in a similar

manner.  One of the ports is selected when the corresponding bit from

port C goes low (PC0-PC3), then the data is read through port B.  PC7 is

the signal SELF which is a software controlled clock for the analyzer.

All analyzer functions are controlled by the outputs at U10.  S0, S1 and

S2 select the system clock source.  MR* is the master rest signal.  WR*

is the write enable for data acquisition.  DIR is the direction control

for U7 and U8, it remains low in normal (non-diagnostic) operation.

TRIG and STCH control the trigger setup.  The outputs at U9 are the

trigger words.  U7 and U8 read or write (diagnostic modes only) the data

bus.  U5 and U6 read the address bus and some status bits.

The clock generator consists of the crystal oscillator X1 and a single

GAL16V8, U15.  The oscillator is selected to be twice the frequency of

the maximum clock rate (a 50 Mhz oscillator for a 25 Mhz analyzer).  The

GAL16V8 is programmed to act as a 5 bit binary counter and an 8 to 1 MUX

with a two input OR gate thrown in.  The output of the oscillator is

feed into the clock input (pin 1) of U15.  Q0, Q1, Q2, Q3 and Q5 are the

counter outputs.  Q0 is half the input clock frequency, Q1 is a quarter

of the input frequency and so on.  The actual system clock is output on

two pins SC0 and SC1.  The system clock signals SC0 and SC1 are

indentical, they are output on two pins to reduce loading on the output

drivers (the high speed design rule of thumb is to drive a maximum of of

4 to 6 inputs with each output, these outputs each drive 4 inputs).  The

source of the system clock is selected by S0, S1 and S2 as follows:

S2 S1 S0	System clock

1  1  1		Q0 - max rate

1  1  0		Q1 - 1/2

1  0  1		Q2 - 1/4

1  0  0		Q3 - 1/8

0  1  1		Q4 - 1/16

0  1  0		disabled

0  0  1		EXT - external clock

0  0  0		SELF - software clock

The first five options select an internal clock derived from the

oscillator.  The EXT signal is a clock from an external source usually

from the system under test.  The SELF signal is a software generated

clock from the host PC, this allows the software to control the clock

for retrieving data, control and diagnostic functions.

The data storage subsystem consists of U1, U2, U3, U4, U11 and U12.  U1

and U2 are 74ACT574 eight bit latches.  During data acquisition the

outputs of the latches are enabled and data is strobed into the latches

from the probe assembly on the rising edge of the system clock.  The

data appears on the outputs of U1 and U2 within 10 ns.  The system clock

is connected to the CS* inputs of the 8K x 8 static RAMs (U3 and U4).

When acquiring data the WR* input of the RAMs is always asserted,

therefore data is written to the RAMs on the low half of the system

clock cycle.  The RAM address generation is handled by U11 and U12,

these two chips combine to form a 13 bit counter with some special

control and status features.  The 13 bits of the counter directly drive

the 13 address lines of the static RAMs.  The WRAP output is a status

bit that is set when the counter wraps around from 1FFF to 0.  Once set,

WRAP remains set until a reset occurs.  U11 and U12 are reset on the

rising edge of the system clock when MR* is asserted low, this clears

WRAP and zeros the address bits.  The counting is controlled by the CNT

signal from U14, as long as CNT is high the address counting continues.

The trigger control subsystem consists of U13, U14, U16 and U17.  U16

and U17 check the lower eight bits of the data bus for a match to the

desired trigger condition.  The trigger condition is expressed in two

eight bit words, the trigger word and the "don't care" word.  The test

for a trigger condition is performed on each of the data bits in

parallel.  Looking at just bit 0 we can see how this test works.  If X0

("don't care" word bit 0) is 0, then it is a valid trigger (for that

bit) when T0 (trigger word bit 0) is the same as D0.  If X0 is 1 and T0

is 1 then bit 0 is a "don't care" bit and the trigger for that bit is

always valid regardless of the state of D0.  If X0 is 1 and T0 is 0 then

this is a never trigger state.  The never trigger is used by the

software to disable the trigger logic.  When all 8 bits have a valid

trigger condition then a valid trigger may occur.  U16 and U17 are

identical dual 2 bit comparators that support the don't care and never

trigger conditions.  U16 and U17 will latch the trigger word so that the

XD bus can be multiplexed to handle both the trigger and "don't care"

words.  To set up a trigger condition the trigger word is output to U9

and appears on the XD bus.  On the rising edge of the TRIG signal (from

U10) the data is latched onto the registers T0, T1, T2 and T3 on U16 and

U17.  The trigger word is now stored in the comparators.  The "don't

care" word is then output through U9 onto the XD bus.  The comparators

are now ready for operation.  U13 and U14 handle posttrigger countdown.

In this logic analyzer data acquisition will continue for 4096 clock

cycles after a valid trigger has occurred.  For U13 to respond to a

trigger state its VAL input must be high.  VAL comes from the address

counter and will go high and stay after a minimum address value is

reached, this prevents an accidental trigger during initialization.

Once the VAL signal is high a trigger condition occurs when the inputs

TRG0, TRG1, TRG2 and TRG3 (from U16 and U17) are all low at the rising

edge of the system clock.  If the STCH input is low, output TRG will go

high indicating a trigger has occurred and the countdown will start.  If

STCH is high TRG0, TRG1, TRG2 and TRG3 must remain low for two

consecutive rising system clock edges for TRG to go active and the

countdown to start.  Once the countdown has started, U13 and U14 will

count for 4096 cycles and then bring CNT low.  Bringing CNT low signals

the end of data acquisition and stops both the trigger countdown and the

address counters (U11 and U12).  The PC software on seeing CNT low

begins retrieving the acquired data from the RAMs.  U13 and U14 are

reset by MR*, this reactivates CNT.  When data is to be retrieved by the

host PC U16 and U17 are programmed to never trigger, U1 and U2 are

disabled, WR* is made inactive, the counters are reset making CNT active

and the address 0.  S0, S1 and S2 are brought low making SELF the clock

source.  The data is read through U7 and U8 when SELF is low.  The

address increments when SELF goes high.  In this manner the data is

retrieved into the PC.

BUILDING THE LOGIC ANALYZER

Before building the logic analyzer get a copy of the software from one

of the sources in the parts list.  You will need the software for

testing the analyzer.  The program doubles as a demo - you can see the

capabilities before actually building the analyzer.

The interface board is best built on a purchased PC board (see parts

list for details).  If you elect to hand wire on a PC/XT prototype board

use a heavy gauge wire for the +5 volt and ground connections from the

PC bus to P1.

The main board is best built on the purchased PC board (see parts list

for details).  Other then the installation of the QPACs board assembly

is straight forward.  The three QPACs are installed under each row of

ICs.  The first QPAC is placed in the row with U10, U7, U15, U2, U1 and

U16.  The QPAC pins should go into pin 10 (ground) and pin 20 (power) of

each chip, the one extra pin sits in the via at label Q1 (where it is

tied to the +5 volt plane).  Solder the pin at Q1 first to hold the QPAC

in place and then solder the remaining chips or sockets in the row.  The

chips or sockets will sit on top of the QPAC.  The other two rows are

handled in the same way.  It is a good idea to use sockets on U1 and U2.

You should also use sockets on U3, U4, U11, U12, U13, U14, U15, U16, U17

and X1 if you would like to be able to upgrade the logic analyzer clock

speed in the future.  A socket for X1 would allow the change to a slower

clock if desired.  U3 and U4 require 28 pin skinny DIP sockets, these

can be made by putting two 14 pin DIP sockets back to back.  You can use

sockets for the remaining chips if desired.  The analyzer can be put

into any suitable case.

THE PROBE ASSEMBLY

Believe it or not the probe assembly was the one of design areas that

was toughest to keep the cost down.  Most logic analyzers have a main

unit and ribbon cables that go out to pods.  Each pod has a group of

short wires with clips on the ends.  Typically there is a pod for every

eight data bits.  This type of setup quickly gets rather expensive, so

for the low cost analyzer we use a build it yourself approach.  The main

analyzer unit actually acts as our pod, we only need to add wires and

clips for each data bit (and the external clock and ground reference).

Since we will need 18 clips for a 16 channel unit (10 for an 8 channel

unit) the probe assembly still ends up as a rather expensive part of the

analyzer.  Usable clips run from about $0.65 to $1.50 each.  The more

expensive minipincer clips are easier to work with.  The clips will

attach to the end of a 20 ribbon cable with an IDS connector.  The cable

should not be longer then about 18 inches.  Split pins 1 and 2 off from

the rest of the cable.  Solder both pins 1 and 2 onto a clip and label

this clip CLOCK.  Next spilt 3 and 4 and solder them into a clip, this

is labeled GROUND.  Now starting from pin 20 separate each wire from the

cable and solder into a clip.  The labels go from 0 to 15 (pin 20 is

labeled 0).  The cable should be split most of the way down the

connector.  If the cable is not split you will get crosstalk between the

adjacent wires.  Of course the wires can become a real mess (just like

an expensive analyzer), I haven't found a good way to keep them neat

without bundling them together and getting crosstalk.

BRINGING THE LOGIC ANALYZER UP

The interface board should be tested before hooking up the analyzer

board, it will be easier to locate and fix interface problems isolated

from the main board.  Before plugging the interface into the PC

carefully check for shorts and other errors.  Set the I/O address using

the DIP switch.  I've found that 390H is usually free, this corresponds

to A9-OFF, A8-OFF, A7-OFF, A6-ON, A5-ON, A4-OFF, A3-ON (remember ON

corresponds to a low because it grounds the input).  Once set plug in

the board and power up the PC, if it does not boot up normally, power

down and recheck everything.  The configuration file RGBLOGIC.CFG should

be edited to match the logic analyzer.  RGBLOGIC.CFG is a four line

ASCII text file, the first line contains the base I/O address in hex

(i.e.  390), the second line should always have 1, the third line is the

oscillator (X1) frequency in Mhz (40,50,60) and the fourth line is the

number of channels (8,16).  Any characters after the numeric value on

each line are ignored.  Edit the file to match your configuration.

Check the voltages at the end of the 50 conductor ribbon cable for

proper power and ground supplys.  Start up the diagnostic program

RGBDIAG, select interface test (F7) from the menu.  The interface test

allows you to read and write the 8255 registers.  Write an 80 (all

values are expressed in hex) to the CNTRL register, this will program

ports A, B and C as output ports.  Check each bit of the three ports at

the end of the cable.  You can write different values to the ports from

the diagnostic program.  If the interface doesn't respond at all make

sure the configuration file and DIP switch settings match, if they do

then check the interface board for errors.  If everything is working,

power down.

Double check the analyzer board for shorts or other errors.  Connect the

50 conductor ribbon cable to the analyzer board and power up the PC.  If

the PC does not boot up normally the power supply is probably shorted -

power down and correct the problem.  Check the power supply voltage at

the labeled test points on the analyzer board it should be between 4.75

and 5.25 volts.  Start the diagnostic program RGBDIAG.  Press F1 to

perform the reset test, this will reset the address and trigger

counters, then display the value read from U7 and U8.  It should read

8000, CNT high, WRAP, TRG and A0-A13 low.  If this or any of the next

four tests do not work correctly skip to the low level diagnostics

below.  Next press F2 to perform the address counting test.  This test

will do a reset, then use the SELF clock to increment the address

counter, checking for a proper address for each clock cycle.  Any errors

will be reported.  Next press F3 to perform the clock test.  This test

does a reset, selects an internal clock and reads the address counter.

It does this for each clock, max rate, 1/2, 1/4, 1/8 and 1/16.  The

address count is printed each time.  The actual value is not important

but each value should be about half the previous value.  Next press F4

for the memory test.  This will do a quick test of the memory, any

errors will be reported.  Next press F5 for the trigger test.  This will

test the trigger circuit, any errors will be reported.  These test do

not fully test the analyzer, they are not run at full speed and don't

thoroughly test each chip, but they should catch 99% of assembly errors.

If tests passed you may skip the next section.

If the analyzer failed any of the previous tests you will need to use

the low level diagnostics (F6) and a logic probe (or DVM) to isolate the

problem.  The low level diagnostics allow you to control individual

signals and read through the input ports.  You can switch signals and

monitor them with the logic probe to make sure the transitions occur or

read the input ports and verify that each bit was read correctly.  If

the reset test failed check MR*, SELF, S0, S1, S2, SC0 and SC1 at at

each chip for proper operation.  If the address counting test failed

check (in addition to the above) CNT, A0-A12, PC2, PC3 U5-1 and U6-1.

If the clock test failed also check Q0, Q1, Q2, Q3 and Q4 on U15 (Note -

the logic probe must be fast enough to detect the high speed clocks i.e.

a 10 Mhz logic probe will not detect a 20 Mhz clock).  If the memory

test failed check DIR, WR*, U7-1, U8-1, PC0, PC1, U1-1, U2-1, RD*,

D0-D15, U3-26 and U4-26.  If the trigger test failed check TRIG,

XD0-XD7, STCH, TRG0, TRG1, TRG2 and TRG3 on U13 and T0, T1, T2 and T3 on

U16 and U17 and D0-D7.  Correct any problems found and make sure all

diagnostic tests pass before proceeding.

USING THE LOGIC ANALYZER

At this point you have passed through all of the diagnostics and we're

ready to go, but first a couple of things to beware of.  Always start up

the logic analyzer before connecting the probes to the circuit under

test, I've found that having the analyzer powered off and connected to a

powered circuit can burn up U1 and U2.  More important - before hooking

up to a new test circuit measure check to be sure that there is no AC

potential between the grounds of the circuit and the logic analyzer.

That aside for a first run lets use the analyzer on itself.  Startup the

program RGBLOGIC, the main menu should appear.  Connect probe 0 to A6

(U12-19), probe 1 to A7 (U12-18), probe 2 to A8 (U12-17), probe 3 to A9

(U12-16), probe 4 to A10 (U12-15), probe 5 to A11 (U12-14), probe 6 to

A12 (U12-13) and probe 7 to WRAP (U12-12).  We do not need to connect

the GROUND or CLOCK probes.  Using menu choice F1 setup the trigger to

be 1XXX XXXX, meaning that we want to trigger when probe 7 is high (in

this case WRAP is 1) and we don't care about the state of the other

bits.  F2 will select whether the trigger must be valid for one or two

consecutive clock cycles.  Select one for now.  Use F3 to select the

highest clock rate.  Press F4 to begin data acquisition.  Within a few

seconds the acquisition will be complete and the program will return to

the main menu.  Select F5, this will show the acquired data in STATE

format.  For each clock cycle the state display shows the clock cycle

number relative to the trigger point.  The trigger point is considered

clock cycle zero, the negative numbers are the clock cycles before the

trigger and positive are after the trigger.  The next number is the time

in micro seconds relative to the trigger.  Then the data is displayed

first in hex and then in binary (probe 15 is the most significant bit,

probe 0 the least).  At the trigger point we should see 0 0.000 xx80

xxxx xxxx 1000 0000 (the x is used to represent an unknown value because

we did not connect probes 8 - 15).  You can move forward or backward

through the state display by pressing Pg Up, Pg Dn, Home, End and the

numeric pad 5 key (this may only work with Num Lock active depending on

your keyboard).  Pg Up and Pg Dn will move backward or forward one page

at a time.  Home will jump to the first data acquired and End to the

last data acquired.  The 5 key will return you back to the trigger

point.  A special feature is the search capabilities, by pressing S you

will see the search menu.  Specify the search pattern in the same manner

as selecting a trigger condition (1,0 and X), press return when

finished.  Select F (forward search) or R (reverse search).  The program

will search for data that matches your condition and display that screen

of data.  This is a very powerful feature because of the rather deep 8K

data buffer (many expensive logic analyzers have only 2K deep buffers).

Press ESC to exit the state display.  Press F6 to enter the timing

display (Note - this requires a CGA or compatible display (EGA or VGA)).

The data will be displayed in timing diagram format.  There is a cursor

at the trigger point and the state information from the cursor position

is displayed at the top of the screen.  The ESC, Pg Up, Pg Dn, Home,

End, numeric 5 key and the search option work the same way as with the

state display.  In addition the left and right arrow keys can move the

cursor back and forth through the data to display the state information

for any point on display.  F7 the file option on the menu allows you the

save and restore data and setup information (trigger word, clock rate

etc.).  You can use this to save data and view it at another time.

WRAPPING UP

That covers the details of a low cost logic analyzer.  As the faster ICs

come down in price it may be possible to upgrade the speed from the 25

Mhz current top rate.  Right now the analzer is fast enough for most

microcontroller and 8 or 16 bit microprocessor projects.